MARYLAND DEPARTMENT OF THE ENVIRONMENT 1800 Washington Boulevard, Suite 750 • Baltimore Maryland 21230 1-800-633-6101 ext. 3193 • www.mde.maryland.gov preventive.maintenance@maryland.gov

RADIOLOGICAL HEALTH PROGRAM

INTRAORAL DENTAL/VETERINARY DENTAL PREVENTIVE MAINTENANCE REPORT

| FACILITY NAME: |  |  |  |  |  |  |  |  |  |  | FACILITY CONTACT NAME: |  |  | CONTACT TELEPHONE NO.: |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| FACILITY REGISTRATION NO.: |  |  | Service Provider |  |  |  |  | REGISTERED SERVICE |  |  |  |
| - |  |  |  |  |  |  |  |  |  |  |  |  | Meter Manufacturer: |  |  |  |  | PROVIDER NAME: |  |  |  |
| MDE MACHINE NO. AND SUFFIX: |  |  | Meter Used – |  |  |  |  |  |  |  | Service Provider |  |  |  |  |
|  | / |  |  |  |  |  |  |  |  |  | Model: |  |  |  |  |  |  | Registration Number: |  |  |  |
| Component Use: |  | Film Speed: |  |  | Model Number: |  |  |  |  |  |  | NAME OF SERVICE |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | PROVIDER: |  |  |  |
| Machine Manufacturer: |  |  |  |  |  |  |  |  |  | Calibration Date: |  |  |  |  |  | DATE OF SERVICE: |  |  |  |
| Facility-Designated Room |  |  |  |  |  |  |  | Note any corrective | services |  |  | Does Radiation Machine |  |  |  |  |
| Number: |  |  |  |  |  |  |  |  |  |  |  |  | provided: |  |  | Meet PM Requirements? |  |  |  |
| Tube Serial Number: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Date Facility Owner Made |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Aware of Service Findings: |  |  |  |
| Other information on tube serviced (optional) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Date Corrective |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Action Taken: |  |  |  |
|  | As Found Settings |  |  | Test Settings |  |  |  |  |  |  |  | Preventive Maintenance Data |  |
| KVP |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | PM Interval (Months) |  | 6____ | 12____ 24____ 36____ | 48 ____ |
|  |  |  |  |  |  |  |  |  |  |  | KVP |  |  |  |  |  |  |  |  |  |  |
| mA |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Next PM Due (Date) |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  | mA |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Time: _________ mSec _________ Pulses |  |  |  |  |  |  |  |  | Notes: |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  | Time |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| HVL |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | TESTING |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | X-ray Tube Voltage |  | Min. HVL |
|  |  |  | KVP |  | Timer |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Designed |  | Measured |  | Specified |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Operating |  | Operating |  | Dental |
| Exp 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Range |  |  | Range |  | Systems |
| Exp 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Below 51 |  |  | 30 |  |  |  | 1.5 |
| Exp 3 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 40 |  |  |  | 1.5 |
| Avg |  | 0.00 |  | 0.00 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 50 |  |  |  | 1.5 |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 51–70 |  |  | 51 |  |  |  | 1.5 |
| % Diff |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 60 |  |  |  | 1.5 |
| Mfr. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 70 |  |  |  | 1.5 |
| Spec |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Above 70 |  | 71 |  |  |  | 2.1 |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 80 |  |  |  | 2.3 |
|  |  |  | Test Results |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 90 |  |  |  | 2.5 |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 100 |  |  |  | 2.7 |
|  | Item |  |  |  | Pass |  | Fail |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  | 110 |  |  |  | 3.0 |
| KVP |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 120 |  |  |  | 3.2 |
| Timer |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 130 |  |  |  | 3.5 |
| HVL |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 140 |  |  |  | 3.8 |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 150 |  |  |  | 4.1 |
| Timer Reproducibility |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

By physically and/or electronically signing this report, I attest that this radiation machine is operating within the specifications and guidelines provided by the manufacturer's manual and that the registrant has received a copy of this report for their

| records. |  | Service Provider Initials [ | ] |
| --- | --- | --- | --- |
| Printed Name | Registrant Signature |  | Date |  |
| Printed Name | Service Provider Signature |  | Date |  |

| MDE/ARMA/COM.13 (MDE RX-33) | E-mail report to preventive.maintenance@maryland.gov or fax to 410-537-3198. | Page 1 of 3 |
| --- | --- | --- |
| Form Date 11/14/16 |  | Recycled Paper |
| TTY Users 1-800-201-7165 |  |  |

MARYLAND DEPARTMENT OF THE ENVIRONMENT 1800 Washington Boulevard, Suite 750 • Baltimore Maryland 21230 (410) 537-3193 • 1-800-633-6101 • www.mde.maryland.gov

RADIOLOGICAL HEALTH PROGRAM

Instructions for Intraoral Dental Preventative Maintenance Report

General Information

COMPLETE ONE FORM MDE RX-33 PER TUBE. Completely and legibly fill out the facility information, machine information and service provider information. Include facility room number or name as designated by the facility.

As Found Settings

Record the “as found” setting of the kVp, mA, time and half layer value of the radiation machine. If tested with values other than As Found Settings, document these test values in the block Test Settings.

Preventive Maintenance Data

Record the manufacturer’s recommended preventive maintenance schedule as indicated in the radiation machine manual. If no preventive maintenance schedule exists for the machine, a 12 month maintenance frequency should be used. Record the date of the next scheduled Preventive Maintenance.

Timer Accuracy

For Factory Tolerance-

- Average all exposures.

- Use formula- ((Average time measured – “as found” time)/ “as found” time) X 100 = % of deviation [disregard the sign].

- If the % deviation is within the manufacturer’s recommendation, the unit is in compliance.

- Machine passes or fails with appropriate documentation.

For Uncertified (+/- 10%)-

- Average all exposures.

- Multiply the time set by .10 to get the + or – 10% variable.

- Add the variable to the time set, and then subtract the variable from the time set.

- The two numbers establish the range.

- If the average time measured falls between the two numbers, the machine is in compliance.

kVp Accuracy

For Factory Tolerance-

- Average all exposures.

- Use formula- ((Average kVp measured – “as found” kVp)/”as found” kVp) X 100 = % of deviation [disregard the sign].

- If the % deviation is within the manufacturer’s recommendation, the unit is in compliance.

- Machine passes or fails with appropriate documentation.

For Uncertified (+/- 10 %)-

- Average all exposures.

- Multiply the kVp set by .10 to get the + or – 10% variable.

- Add the variable to the kVp set, and then subtract the variable from the kVp set.

- The two numbers establish the range.

- If the Average kVp measured falls between the two numbers the machine is in compliance.

Timer Reproducibility: T > 5 (Tmax – Tmin)

- Use the timer data from the reverse of this form (Measured and Average).

- Subtract the minimum time from the maximum time (Measured values).

- Multiply the result by the factor of 5 as shown above.

- Compare to the average of the measured values for time.

- If the average of the measured values is greater than or equal to the multiplied result, the timer is reproducible. (PASS)

For Dental Preventive Maintenance Use Only

MDE/ARMA/COM.13 (MDE RX-33)	Page 2 of 3

Form Date 11/14/16	Recycled Paper

TTY Users 1-800-201-7165

MARYLAND DEPARTMENT OF THE ENVIRONMENT 1800 Washington Boulevard, Suite 750 • Baltimore Maryland 21230 (410) 537-3193 • 1-800-633-6101 • www.mde.maryland.gov

|  |  |  |  |  | RADIOLOGICAL HEALTH PROGRAM |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |  | MENU |  |  |  |  |
|  |  | 05. CODE PROFESSION |  |  | CODE DARKROOM |  | 10. | (continued) |  |
|  |  |  |  |  |  |  |  |  |  | 30 | Dynavision |  |
|  | 10 |  | Hospital |  |  | AP | Automatic Processor |  |  |  |
|  | 11 |  | Chiropractor |  |  | DD | Complete Digital Imaging | 31 | E.G. & G. |  |
|  | 12 |  | Dentist |  |  | IP | Insta-fix only processing | 25 | Elekta |  |
|  | 13 |  | Physician |  |  | MP | Manual Processing | 20 | Faxitron |  |
|  | 14 |  | Podiatrist |  |  | NP | No processing on-site | 21 | Fischer Imaging Group |  |
|  | 15 |  | Radiologist |  |  |  |  |  | 34 | Fuji |  |
|  | 16 |  | Industrial/Field Radiography |  |  | CODE MEDICAL THERAPY | 23 | Gendex |  |
|  | 17 |  | Veterinarian |  |  | AT | Accelerator | 24 | General Electric |  |
|  | 18 |  | State/Local Government |  |  |  |  | 35 | Glenbrook |  |
|  | 19 |  | Education/Research |  |  | CT | Contact Therapy | 37 | Global Marine |  |
|  | 20 |  | Portable/Mobile X-ray |  |  | DT | Deep X-ray | 39 | Golden |  |
|  | 21 |  | Other |  |  | ST | Superficial | 40 | HCMI |  |
|  |  |  |  |  |  |  | CODE INDUS/EDUC/RESEARCH | 41 | Heimann |  |
|  |  | 09. COMPONENT USE |  |  |  | 46 | Heuft Systems Technik |  |
|  |  |  |  |  |  |  | IA | Accelerator | 27 | Hewlett-Packard |  |
|  |  | CODE DENTAL |  |  |  |  | 28 | Hitachi |  |
|  |  |  |  |  | IC | Cabinet Radiography |  |  |  |
|  |  |  |  |  |  |  |  |  | 38 | Hologic |  |
|  |  | CBCT Cone Beam Computed Tomography |  |  | IE | Electron Microscope |  |  |  |
|  |  |  |  |  |  |  | 48 | Hope |  |
|  |  | CD | Cephalometric |  |  | IF | Field Radiography | 43 | Instrumentarium |  |
|  |  | CP | Cephalometric/Intra-oral Comb. |  |  | IG | Gauge | 55 | JEOL |  |
|  |  | CX | Pan/Ceph Combination |  |  | IN | Diffraction | 32 | J. Morita |  |
|  |  |  |  |  |  | IO | Other Indus./Educ./Research |  |  |  |
|  |  | HH | Hand-held |  |  |  |  | 33 | Kodak |  |
|  |  | ID | Intra-oral |  |  | IR | Room Radiography | 44 | Konica |  |
|  |  | XD | Panorex |  |  | IS | Spectrographic | 56 | LG |  |
|  |  | TD | TMJ Work |  |  | CODE MEDICAL FLUOROSCOPE | 47 | Lorad |  |
|  |  | OD | Other Dental |  |  |  | 36 | Lumix |  |
|  |  | CODE VETERINARY |  |  | AF | Above Table Tube | 49 | Lunar |  |
|  |  |  |  | BF | Below Table Tube | 50 | Midwest/Sybron |  |
|  |  | VP | Veterinary Portable |  | CF | C-Arm | 57 | Min X-ray |  |
|  |  |  |  |  |  | MF | Mobile Fluoroscope | 61 | Niton |  |
|  |  | VS | Veterinary Stationary |  |  |  |  |  |  |  |
|  |  |  |  |  |  | UF | Upright Fluoroscope | 42 | OEC Diasonics |  |
|  |  | VD Veterinary Dental |  |  |  |  |  |  |  |
|  |  |  |  | OF | Other Medical Fluoroscope | 66 | PANalytical |  |
|  |  | CODE MEDICAL |  |  |  |  |  | 59 | Panoramic Corp. |  |
|  |  |  |  |  | 10. CODE MANUFACTURER | 45 | Phillips |  |
|  |  | AD | Angiography/Digital |  |  | 00 |  | Imagie Works | 60 | Planmeca |  |
|  | AN | Angiography |  |  |  |  |  | 70 | Progeny |  |
|  | BD | Bone Densitometry |  | 01 |  | AS and E | 72 | Protec |  |
|  |  | CA | CAT Scanner |  | 02 |  | Accuray | 74 | Rapiscan |  |
|  |  | CE | Ceiling Tube (Leg Studies) |  | 06 |  | Accudex | 51 | Raytheon |  |
|  |  | CH | Chest, Dedicated |  | 07 |  | Acoma | 73 | Rigaku |  |
|  |  | CI | Chiropractic |  | 03 |  | Agfa | 52 | Ritter |  |
|  |  | DI | Diathermy |  | 08 |  | Air Techniques | 53 | S.S. White |  |
|  |  | GP | General Purpose |  | 14 |  | All Pro | 54 | Sanko |  |
|  |  | HN | Head and Neck |  | 04 |  | Andrex | 78 | Sedecal |  |
|  |  | MA | Mammography |  |  | 85 Aribex | 79 | Seiko |  |
|  |  | MI | Magnetic Imaging |  | 05 |  | Asoma | 58 | Siemens |  |
|  |  | OT | Other Medical |  | 10 |  | Astrophysics | 80 | Sirona |  |
|  |  | PD | Podiatry |  | 12 |  | Autoclear | 64 | Soredex |  |
|  |  | PH | Portable Hand Carried |  | 16 |  | Aztech | 81 | Spectro |  |
|  |  | PM | Portable Mobile |  | 09 |  | Belmont | 68 | Summit |  |
|  |  | SR | Stereotactic |  | 11 |  | Bennett X-ray | 62 | Toshiba |  |
|  |  | TO | Tomography |  | 13 |  | Bowie | 63 | Transworld |  |
|  |  | UR | Urology |  | 18 |  | Castle | 71 | Trophy |  |
|  |  | US | Ultrasound |  | 15 |  | Continental X-ray Corp. | 65 | Universal |  |
|  |  |  |  |  |  | 17 |  | Control Screening | 67 | Varian |  |
|  |  |  |  |  |  | 19 |  | Coromex | 82 | Vet Ray, Inc. |  |
|  |  |  |  |  |  | 26 |  | de Gotzen | 69 | Weber |  |
|  |  |  |  |  |  | 29 |  | Del Medical | 83 | XMA |  |
|  |  |  |  |  |  | 22 |  | Dentx | 84 | X-Cel |  |
|  |  |  |  |  |  |  |  |  |  | 76 | Yoshida |  |
|  |  |  |  |  |  |  |  |  |  | 77 | Other |  |

MDE/ARMA/COM.013 (MDE RX-33)	Page 3 of 3

Form Date 11/14/16	Recycled Paper

TTY Users 1-800-201-7165