Pira Consulting Report

Logistics, Distribution & Operations Group

Demonstration of compliance to UN Packing Instruction 650 for packaging for diagnostic specimens

05K19J0856

Prepared for Julia Smallbone of Versapak

By Justin Scott, Consultant

June 2005

Private and confidential

Pira International Limited

Registered office: Randalls Road, Leatherhead, Surrey, KT22 7RU, United Kingdom

Main line +44 (0)1372 802000 Facsimile +44 (0)1372 802249

Registered number: 3858209 England

Pira International – Versapak – 05K19J0856

Demonstration of compliance to UN Packing Instruction 650 for packaging for diagnostic specimens

- Introduction

This report covers Pira’s laboratory distribution testing of packaging for diagnostic specimens for:

Versapak International Ltd

The Versapak Centre

Centurion Way

Erith

Kent

DA18 4AF

- Background

Versapak wishes to supply heavy duty padded packs for diagnostic specimens. In order to meet carriage of dangerous goods regulations the pack must be capable of meeting certain performance requirements. Versapak approached Pira International to conduct the relevant tests on three pack sizes.

Important note, the regulations make it clear that packs must be ‘capable of passing’ the tests outlined below. It is not a formal requirement of the dangerous goods regulations that these tests be conducted by an independent laboratory.

- Objective

To demonstrate compliance with the drop test requirements of United Nations Packing Instruction 650 for packaging for diagnostic specimens

- Samples

Five packs of each of 3 sizes were sent to Pira for testing along with 1 pack of vials as a typical diagnostic specimen.

A summary of the packaging submitted for testing is presented in Table 1 with photographs of samples in Appendix A.

1

Pira International – Versapak – 05K19J0856

Demonstration of compliance to UN Packing Instruction 650 for packaging for diagnostic specimens

Table 1 Summary specification of packaging submitted

| Description | 3 sizes of heavy duty padded hold-all style packs for |
| --- | --- |
|  | diagnostic specimens |  |
| Sample | Pack 1 | Pack 2 | Pack 3 |
| External | 300x200x300 | 480x280x300 | 660x280x370 |
| dimensions | mm | mm |  |
| Empty weight | 1.63 kg | 2.53 kg | 3.92 kg |
| Manufacturer | Versapack |  |  |
| Method of closure | Zipped |  |  |
| Material | PVC coated nylon with padding |  |

- Work undertaken

Pira filled out the packs with the vials supplied and a range of liquid filled containers as detailed in Table 2

Table 2 Pack contents for testing

|  | Vials in plastic | Liquid filled | Total filled |
| --- | --- | --- | --- |
|  | holder & plastic |  | pack weight |
|  |  | plastic jars |  |
|  | bag |  |  |
| Pack 1 | 1 | 0 | 2.4 kg |
| Pack 2 | 1 | 2 off | 4.96 kg |
| Pack 3 | 1 | 3 off | 7.65 kg |

Pira then conducted testing in accordance with the requirements of Packing Instruction 650 of the United Nations Recommendations on the Transport of Dangerous Goods (13th edition) comprising:

- Drop tests from a height of 1.2m on five complete pack systems (of each size) through the following drops:

  - Flat onto base

  - Flat onto top

  - Flat onto longest side

  - Flat onto shortest side

  - Onto a corner

2

Pira International – Versapak – 05K19J0856

Demonstration of compliance to UN Packing Instruction 650 for packaging for diagnostic specimens

- Results

Packs and contents were inspected following each drop sequence. The results are presented in Table 3-5.

Table 3 Observations following drop testing – Pack 1

| Pack 1 | Drop | Observations |
| --- | --- | --- |
| Sample A | 1 | No damage to pack or contents |
| Sample B | 2 | No damage to pack or contents |
| Sample C | 3 | No damage to pack or contents |
| Sample D | 4 | No damage to pack or contents |
| Sample E | 5 | No damage to pack or contents |
| Table 4 Observations following drop testing – Pack 2 |
| Pack 2 | Drop | Observations |
| Sample A | 1 | No damage to pack or contents |
| Sample B | 2 | No damage to pack or contents |
| Sample C | 3 | No damage to pack or contents |
| Sample D | 4 | No damage to pack or contents |
| Sample E | 5 | No damage to pack or contents |
| Table 5 Observations following drop testing – Pack 3 |
| Pack 3 | Drop | Observations |
| Sample A | 1 | No damage to pack or contents |
| Sample B | 2 | No damage to pack or contents |
| Sample C | 3 | No damage to pack or contents |
| Sample D | 4 | No damage to pack or contents |
| Sample E | 5 | No damage to pack or contents |

3

Pira International – Versapak – 05K19J0856

Demonstration of compliance to UN Packing Instruction 650 for packaging for diagnostic specimens

- Conclusion

The packs provided satisfactory protection to the replicated diagnostic specimens within and are therefore considered compliant with the drop test requirements of United Nations Packing Instruction 650 for packaging for diagnostic specimens.

Prepared by	Checked by

Justin Scott	Khalid Ahmad

Consultant	Consultant

4

Pira International – Versapak – 05K19J0856

Demonstration of compliance to UN Packing Instruction 650 for packaging for diagnostic specimens

Appendix A

Photographs

Photograph 1 Vial pack	Photograph 2 Case 1 and contents

Photograph 3 Case 2 and contents	Photograph 4 Case 3 and contents

Photograph 5 Case 1 zipped up	Photograph 6 Example of drop test

5